## Slide 1: 2-channel Oscilloscope + LabVIEW + Transducers can be used for multiple acoustic non-destructive evaluation techniques
Frequency Sweep
  Sweep through a range of frequencies, record input (Ch1) and output (Ch2) waveforms
    Example: sweep from 20 kHz to 100 kHz at 1 kHz increments using a 5 V sine wave
    Amplitudes / phase shift can be extracted from input/output waveforms for frequency response
Pulse Generation
  Send (Ch1) and receive (Ch2) pulses
    Example: Send and receive 30 pulses at 5 V with a width of 1 s
    Speed of sound measurements, acousto-ultrasonics
Acquisition
  Short time (order of 100s )
    Acoustic emission monitoring
    NOTE: Holdout time between hits is fast (<< 1 second), but unknown
  Long time (order of 100s ms)
    Modal analysis (hitting with modal hammer/artificial source for resonant frequencies)
  NOTE:
    A larger horizontal scale will have lower sampling rates, recording a longer waveform may not always be better depending on the frequencies you expect to measure
    Lockout time after a trigger event << 1 second, but not quantified/coded into the program

## Slide 2: Piezos hooked up to WaveGen, Ch1, and Ch2 of Keysight InfiniiVision 2000X
Oscilloscope
WAVE GEN
Ch1
Ch2
Piezo 1
Piezo 2
3 way BNC
BNC Cable
Back of oscilloscope has Device port which is connected to PC via USB
PC

## Slide 3
FREQUENCY SWEEP
PULSE GENERATION
CH1 INPUT SIGNAL
CH2 RECEIVED SIGNAL
CH1 INPUT SIGNAL
CH2 RECEIVED SIGNAL

## Slide 4
ACOUSTIC EMISSION
MODAL ANALYSIS

## Slide 5: Equipment, software, drivers
Keysight InfiniiVision Oscilloscope
  Programmed using 2000X model
  Program should work for other InfiniiVision models (3000X, 4000X, etc.) but I have not verified this
Useful links for learning more
  https://www.youtube.com/watch?v=YkmUGGwUhYo
  https://www.youtube.com/watch?v=Q4PpfmaSf5I
  https://knowledge.ni.com/KnowledgeArticleDetails?id=kA03q000000x3XpCAI&l=en-US
Software needed:
  LabVIEW 2017, base development
  IO Libraries Suite
    https://www.keysight.com/us/en/lib/software-detail/computer-software/io-libraries-suite-downloads-2175637.html
    This should allow you to connect oscilloscope to PC via Device port and IO Libraries Suite, VISA address is acquired from this panel
  Drivers specific to Keysight Oscilloscopes that let you program easily in LabVIEW
    Driver 1
      https://www.keysight.com/us/en/lib/software-detail/driver/infiniivision-xseries-oscilloscope-ivi-and-matlab-instrument-drivers-2019021.html
    Driver 2
      https://www.keysight.com/us/en/lib/software-detail/driver/infiniivision-xseries-oscilloscope-labview-instrument-drivers-2862255.html
Any transducers should work
  https://www.steminc.com/PZT/en/piezo-disc
From driver page:

## Slide 6: Procedure:
1. With LabVIEW and drivers properly installed, open program
2. Enter information into SETUP
  Important that file name is changed after every test, otherwise will OVERWITE previously saved .txt files in the working directory with same name
  Oscilloscope VISA Address must be changed to your oscilloscope from Keysight IO Suite / Connection Expert
(will be green checkmark when plugged in)

## Slide 7: Procedure:
3. Enter information into OSCILLOSCOPE CONFIGURATION
  Trigger edge source: which channel waveform triggers an event
  Trigger edge level: voltage threshold which triggers an acquisition for trigger edge source
  Channel 1 / 2 Scale: depends on the amplitude of the waveform you expect to measure
  Offset and attenuation generally don’t need to be changed
  Horizontal scale: this determines the length of signal you will acquire
  Acquisition start time: this determines pretrigger length (make sure to put a negative sign)

## Slide 8: Procedure:
4. If running Frequency Sweep or Pulse Generation test, enter parameters
  Frequency Sweep:
    Amplitude: amplitude of the Ch1 waveform, max is 5 V
    Initial Frequency: Initial sine wave frequency
    Stop Frequency: final sine wave frequency
    Freq. Increments: Increment size from initial frequency until stop frequency
    Acq. Per Freq: How many waveforms acquired per frequency
    Delay: time between each acquisition (too fast will crash oscilloscope, about 1 second appears okay)
  Pulse Generation:
    Pulse width: width of square pulse
    Rate: again, too fast may crash oscilloscope

## Slide 9: Procedure:
5. Once all parameters are entered, hit run program
  Oscilloscope will REBOOT (you will probably hear a clicking sound), and configure itself to the entered parameters in OSCILILOSCOPE CONFIGURATION, but will not start test
    Will begin LOG with entering configuration information and time
  When ready, hit START button for selected test
  Frequency Sweep and Pulse Generation will run to completion, when program completed it will output .txt file of acquired waveforms
  If running Acquisition, end the test by pressing STOP button (don’t press stop sign at the top of the program window, will not save)

## Slide 10: Procedure:
6. Text .txt file will be saved to directory with settings and data
7. Waveform data is saved Ch1 and Ch2, 50000 sample points per waveform
  Total number of data points: 50,000 * number of events
8. Use python script, ‘load_example_outputs.py’ to load data
  If run code directly (__main__ func), it will load in .txt files from example outputs for the 3 type of tests
Data points stored in two long columns

## Slide 11: Troubleshooting / Advice
Find correct configuration for test and oscilloscope generally requires trial and error. But this can all be done MANUALLY using oscilloscope controls while operating in SINGLE mode and using the WaveGen.
If using an AMPLIFIER, be careful at start up: When the scope resets, it will default to a 1V sine wave which when amplified could be much larger in your circuit
CHANGE FILE NAME! Files will OVERWRITE if name is kept the same. DOES NOT APPEND!
Oscilloscope randomly crashes during pulse generation and frequency sweep; it is unclear why this occurs but probably due to rate at which events are occurring.
  Pulse generation “rate” and frequency sweep “delay” can be increased
  Recommendation is to go for shorter tests, and more data files; for instance, perform multiple frequency sweeps (100-200 kHz, 200-300 kHz, etc.) rather than 100-500 kHz.
  AE monitoring for long term mechanical tests not recommended
Sampling rate depends on the length of acquisition
  A larger horizontal scale will have lower sampling rates, so recording a longer waveform may not always be better depending on the frequencies you expect to measure, the sampling rate appears on the oscillscope
Lockout time after a trigger event << 1 second, but not quantified/coded into the program

## Slide 12: Python File Addition – May 3, 2024 ‘store_infiniivision_oscilloscope_data_as_binary’
Modified ‘brute-force’ script from InfiniiVision page
IO Suite Libraries is needed on host computer.
PyVisa library needed in python
Additional Driver for VISA/USB? Might be okay with just IO Suite Libraries (Not sure yet)